## TE-5000
Site Information
Location: | Cleves, Ohio |  | Site ID: | 145 |  | Date: | 2014-10-31
Sampler: | TE-5000 TSP |  | Serial No: | 467 |  | Tech: | Jim Tisch
Site Conditions
Barometric Pressure (in Hg): |  |  | 29.5 | Corrected Pressure (mm Hg): |  |  | 749.3
Temperature (deg F): |  |  | 68 | Temperature (deg K): |  |  | 293.15
Average Press. (in Hg): |  |  | 29.4 | Corrected Average (mm Hg): |  |  | 746.76
| Average Temp (Deg F): |  | 65 | Average Temp: (Deg K): |  |  | 291.4833
Calibration Orifice
| Make: | Tisch |  | Qstd Slope: |  | 1.583
| Model: | TE-5028A |  | Qstd Intercept: |  | -0.0152
| Serial#: | 1179 | Calibration Due Date |  |  | 2015-10-24
Calibration Information
Plate or | H2O | Qstd | I | IC
Test # | (in) | (m3/min) | (chart) | (corrected) |  | Linear Regression
1 | 7.25 | 1.7124 | 52 | 52.058 |  | Slope: | 28.7714
2 | 5.75 | 1.526 | 47 | 47.0524 | Intercept: |  | 2.9455
3 | 4.55 | 1.3586 | 42 | 42.0469 | Corr. Coeff: |  | 0.99984
4 | 3.45 | 1.1842 | 37 | 37.0413
5 | 2.7 | 1.0487 | 33 | 33.0368 | # of Observations: |  | 5
Calculations
Qstd = 1/m[Sqrt(H2O(Pa/Pstd)(Tstd/Ta))-b] |  |  |  |  | m = sampler slope
IC =I[Sqrt(Pa/Pstd)(Tstd/Ta)] |  |  |  |  | b = sampler intercept
I = chart response
Qstd = standard flow rate |  |  |  |  | Tav = daily average temperature
IC = corrected chart response |  |  |  |  | Pav = daily average pressure
I = actual chart response
m = calibrator Qstd slope
b = calibrator Qstd intercept |  |  |  |  | Enter Average I (chart): |  | 38
Ta = actual temperature during calibration (deg K) |  |  |  |  | Average Flow Calculation m3/min
Pa = actual pressure during calibration (mm Hg) |  |  |  |  | 1.2214
Tstd = 298 deg K |  |  |  |  | Average Flow Calculation in cfm
Pstd = 760 mm Hg |  |  |  |  | 43.1278
For subsequent calculation of sampler flow: |  |  |  |  | Sample Time (Hrs): |  | 24
1/m((I)[Sqrt(298/Tav)(Pav/760)]-b) |  |  |  |  | Total flow in 24 hours m3/min
1758.7855
Total flow in 24 hours cfm
62104.0082
NOTE: Ensure calibration orifice has been certified within 12 months of use
Tisch Environmental 145 South Miami Ave, Cleves OH 45002 ● 877.263.7610 ● sales@tisch-env.com ● www.tisch-env.com
TE-5000 TSP Sampler Calibration (Dickson Recorder)