## TE-5000
Site Information
Location: | Cleves, Ohio |  | Site ID: | 145 |  | Date: | 2014-10-31
Sampler: | TE-5170 MFC |  | Serial No: | 367 |  | Tech: | Jim Tisch
Site Conditions
Barometric Pressure (in Hg): |  |  | 29.5 | Corrected Pressure (mm Hg): |  |  | 749.3
Temperature (deg F): |  |  | 68 | Temperature (deg K): |  |  | 293
Average Press. (in Hg): |  |  | 29.65 | Corrected Average (mm Hg): |  |  | 753.11
Average Temp. (deg F): |  |  | 70 | Average Temp. (deg K): |  |  | 294.1111
Calibration Orifice
| Make: | Tisch |  | Qstd Slope: |  | 1.583
| Model: | TE-5028A |  | Qstd Intercept: |  | -0.0152
| Serial#: | 1179 |  | Date Certified: |  | 2014-10-24
Calibration Information
Plate or | H2O | Qstd | I | IC
Test # | (in) | (m3/min) | (chart) | (corrected) |  | Linear Regression
1 | 7.25 | 1.7128 | 52 | 52.0713 |  | Slope: | 28.7714
2 | 5.75 | 1.5264 | 47 | 47.0645 | Intercept: |  | 2.9463
3 | 4.55 | 1.3589 | 42 | 42.0576 | Corr. Coeff: |  | 0.99984
4 | 3.45 | 1.1845 | 37 | 37.0508
5 | 2.7 | 1.049 | 33 | 33.0453 | # of Observations: |  | 5
Calculations
Qstd = 1/m[Sqrt(H2O(Pa/Pstd)(Tstd/Ta))-b] |  |  |  |  | m = sampler slope
IC = I[Sqrt(Pa/Pstd)(Tstd/Ta)] |  |  |  |  | b = sampler intercept
I = chart response
Qstd = standard flow rate |  |  |  |  | Tav = daily average temperature
IC = corrected chart response |  |  |  |  | Pav = daily average pressure
I = actual chart response
m = calibrator Qstd slope |  |  |  |  | Average I (chart): |  | 44
b = calibrator Qstd intercept |  |  |  |  | Average Flow Calculation m3/min
Ta = actual temperature during calibration (deg K) |  |  |  |  | 1.43
Pa = actual pressure during calibration (mm Hg) |  |  |  |  | Average Flow Calculation in CFM
Tstd = 298 deg K |  |  |  |  | 50.4924
Pstd = 760 mm Hg |  |  |  |  | Sample Time (Hrs): |  | 24
For subsequent calculation of sampler flow: |  |  |  |  | Total Flow in m3/min
1/m((I)[Sqrt(298/Tav)(Pav/760)]-b) |  |  |  |  | 2059.1646
Total Flow in CFM
72709.1023
NOTE: Ensure calibration orifice has been certified within 12 months of use
Tisch Environmental 145 South Miami Ave, Cleves OH 45002 ● 877.263.7610 ● sales@tisch-env.com ● www.tisch-env.com
TE-5170 Calibration Worksheet