## TE-6070
Site Information
Location: |  | Cleves Ohio | Site ID: | 145 | Date: | 2014-10-31
| Sampler: | TE-6070 PM10 | Serial No: | 654 | Tech: | Jim Tisch
Site Conditions
Barometric Pressure (in Hg): |  |  | 29.5 | Corrected Pressure (mm Hg): |  |  | 749.3
Temperature (deg F): |  |  | 68 | Temperature (deg K): |  |  | 293
Average Press. (in Hg): |  |  | 29.58 | Corrected Average (mm Hg): |  |  | 751.332
Average Temp. (deg F): |  |  | 66 | Average Temp. (deg K): |  |  | 291.8889
Calibration Orifice
Make: |  | Tisch Environmental, Inc. |  |  | Slope: |  | 0.99128
Model: |  | TE-5028A |  |  | Intercept: |  | -0.00941
Serial#: |  | 1179 |  |  | Calibration Due Date: |  | 2015-10-24
Calibration Data
Plate or | In H2O | Qa | I | IC
Test # |  | (m3/min) | (chart) | (corrected) |  | Linear Regression
1 | 5.75 | 1.5222 | 47 | 29.3903 | Slope |  | 18.4939
2 | 3.8 | 1.2392 | 39 | 24.3877 | Intercept |  | 1.3134
3 | 3.45 | 1.1812 | 37 | 23.137 | Corr. Coeff |  | 0.99981
4 | 2.7 | 1.046 | 33 | 20.6357 |  | SFR | 1.1374
5 | 1.75 | 0.844 | 27 | 16.8838 |  | SSP | 35.7381
[5 empty] | # of Observations: |  | 5
Calculations
Qa = 1/m(Sqrt((H2O)(Ta/Pa))-b) |  |  | SFR = 1.13(Ps/Pa)(Ta/Ts) |  |  | m = sampler slope
IC = I(Sqrt(Ta/Pa)) |  |  | SSP = (m*SFR+b)(Sqrt(Pa/Ta)) |  |  | b = sampler intercept
I = chart response
Qa = actual flow rate |  |  | SFR = sampler set point flow rate |  |  | Tav = daily average temperature
IC = corrected chart response |  |  | SSP = sampler chart set point |  |  | Pav = daily average pressure
m = calibrator slope |  |  | m = sampler slope
b = calibrator intercept |  |  | b = sampler intercept |  |  | Average I(chart): | 36
Ta = actual temperature (deg K) |  |  | Ta = actual temperature (deg K) |  |  | Average Flow over Sample (m3/min)
Pa = actual pressure (mm Hg) |  |  | Pa = actual pressure (mm Hg) |  |  | 1.1423
For subsequent calculation |  |  | Ts = Average temperature (deg K) |  |  | Enter Total Time (Hrs): | 23.9
of sampler flow: 1/m((I)(Sqrt(Tav/Pav))-b) |  |  | Ps = Average pressure (mm Hg) |  |  | Total flow over sample (m3/min)
1638.0278
NOTE: Ensure calibration orifice has been certified within 12 months of use | [5 empty] | Total flow over sample (CFM)
57838.7599
Tisch Environmental 145 South Miami Ave, Cleves OH 45002 ● 877.263.7610 ● sales@tisch-env.com ● www.tisch-env.com
TE-6070 Calibration Worksheet