## TE-5000
Site Information
Location: | Cleves, Ohio |  | Site ID: | 145 |  | Date: | 2018-01-17
Sampler: | TE-5170 MFC |  | Serial No: | 367 |  | Tech: | Bob Tisch
Site Conditions (From TE-HVC Calibrator)
Barometric Pressure (mmHg): |  |  | 29.5 | Temperature (deg C): |  |  | 23.2
TE-HVC Digital Calibrator Information
| Make: | Tisch |  | Date Certified: |  | 2017-10-10
| Model: | TE-HVC |  | Serial#: |  | 12
Calibration Information
Test # | I | Qstd | Qact
| (Chart) | (TE-HVC) | (TE-HVC) |  |  | Linear Regression
1 | 52 | 50.4 | 49.9 |  |  | Slope: | 1.0185
2 | 46 | 44.2 | 43.6 |  | Intercept: |  | -3.2972
3 | 43 | 40.8 | 40.5 |  | Corr. Coeff: |  | 0.99814
4 | 38 | 34.9 | 34.6
5 | 31 | 29.1 | 28.8 |  | # of Observations: |  | 5
Calculations
[5 empty] | Average I (chart): |  | 44
Average Flow Calculation CFM
Qact = actual flow rate from calibrator |  |  |  |  | 41.517
Qstd = standard flow rate from Calibrator |  |  |  |  | Average Flow Calculation in m3/min
1.1625
I = actual chart response |  |  |  |  | Sample Time (Hrs): |  | 24
Slope = slope of chart versus Qact linear regression |  |  |  |  | Total Flow in CFM
Intercept = intercept of chart versus Qact linear regression |  |  |  |  | 59784.4913
Total Flow in m3/min
For subsequent calculation of sampler flow: |  |  |  |  | 1673.9658
Qact = Slope x Chart reading + Intercept (y=mx+b)
NOTE: Ensure TE-HVC digital calibrator has been certified within 12 months of use
Tisch Environmental 145 South Miami Ave, Cleves OH 45002 ● 877.263.7610 ● sales@tisch-env.com ● www.tisch-env.com
MFC Calibration Worksheet Using the TE-HVC Digital Calibrator