## TE-1000
Site Information
| Location: | Cleves Ohio | Site ID: | 145 | Date: | 2014-10-31
| Sampler: | TE-1000 | Serial No: | 1116 | Tech: | Jim Tisch
Site Conditions
Barometric Pressure (in Hg): |  |  | 29.8 | Corrected Pressure (mm Hg): |  |  | 756.92
Temperature (deg F): |  |  | 68 | Temperature (deg K): |  |  | 293.15
Average Pressure (in Hg): |  |  | 29.58 | Corrected Average Pressure (mm Hg): |  |  | 751.332
Average Temperature (deg F): |  |  | 68 | Average Temperature (deg K): |  |  | 293.15
Calibration Orifice
| Make: | Tisch |  |  |  | Qstd Slope: | 9.7769
| Model: | TE-5040A |  |  |  | Qstd Intercept: | -0.04181
| Serial#: | 1185 |  |  |  | Calibration Due Date: | 2014-10-15
Calibration Information
Plate or | Pressure | Qstd | Flow | Flow
Test # | (in H20) | (m3/min) | (magn) | (corrected) |  | Linear Regression
1 | 7.5 | 0.28612 | 70 | 8.4184 | Slope: |  | 35.1848
2 | 6.6 | 0.26867 | 60 | 7.7939 | Intercept: |  | -1.6866
3 | 5.75 | 0.25106 | 50 | 7.1149 | Corr. Coeff: |  | 0.99778
4 | 4.9 | 0.23209 | 40 | 6.3637
5 | 3.7 | 0.20224 | 30 | 5.5111 | # of Observations: |  | 5
Calculations
Qstd = 1/m[Sqrt((H20)(Pa/760)(298/Ta))-b] |  |  |  | m = sampler slope
Flow (corrected)=Sqrt((magn)(Pa/Pstd)(Tstd/Ta)) |  |  |  | b = sampler intercept
(magn) = magnehelic reading
Qstd = standard flow rate |  |  |  | Tav = daily average temperature
Flow (magn)= reading from magnehelic gauge |  |  |  | Pav = daily average pressure
Flow (corrected)= corrected flow rate
m = calibrator Qstd slope | [5 empty] | Average Flow (magn): | 36
b = calibrator Qstd intercept | [5 empty] | Average Flow Over Sample (m3/min)
Ta = actual temperature during calibration (deg K) | [5 empty] | 0.21889
Pa = actual pressure during calibration (mm Hg) | [5 empty] | Enter Total Time (hrs): | 23.9
Tstd = 298 deg K | [5 empty] | Total Flow Over Sample (m3)
Pstd = 760 mm Hg | [5 empty] | 313.8827
For subsequent calculation of sampler flow: | [5 empty] | Total Flow Over Sample (liters)
Qstd = 1/m[Sqrt((H20)(Pa/760)(298/Ta))-b] | [5 empty] | 313882.7485
NOTE: Ensure calibration orifice has been certified within 12 months of use
Tisch Environmental 145 South Miami Ave, Cleves OH 45002 ● 877.263.7610 ● sales@tisch-env.com ● www.tisch-env.com
TE-1000 PUF Calibration Worksheet

## Revison History (hidden)
Revision History: BT - 12-26-13 updated logo and format to match current worksheets added x.x decimal place for condition information for temp and pressure added kelvin temperature calculation to 273.15 instead of 273 Deleted rows 16 and 17 - avg press/temp /corrected seasonal / average temp