## TE-6170V
Site Information
| Location: | Cleves Ohio | Site ID: | 145 | Date: | 2014-10-31
| Sampler: | TE-6070V | Serial No: | P6862 PM10 | Tech: | Jim Tisch
Site Conditions
Temp (deg F): |  | 68
Ta (deg K): |  | 293 |  | Barometric Press (in Hg): |  | 29.5
Ta (deg C): |  | 20 |  | Pa (mm Hg): |  | 749.3
Calibration Orifice
| Make: | Tisch |  | Qa Slope: |  | 0.99128
| Model: | TE-5028A |  | Qa Intercept: |  | -0.00941
| Serial#: | 1179 |  | Calibration Due Date: |  | 2015-10-24
Calibration Information
Run | Orifice | Qa | Sampler | Pf |  | Look Up | % of
Number | "H2O | m3/min | "H2O | mm Hg | Po/Pa | m3/min | Diff
1 | 3.35 | 1.1641 | 5.55 | 10.3578 | 0.98618 | 1.204 | 3.4361
2 | 3.35 | 1.1641 | 6.25 | 11.6642 | 0.98443 | 1.202 | 3.2643
3 | 3.3 | 1.1554 | 7.1 | 13.2506 | 0.98232 | 1.199 | 3.8081
4 | 3.3 | 1.1554 | 7.8 | 14.5569 | 0.98057 | 1.198 | 3.7215
5 | 3.25 | 1.1467 | 9.4 | 17.543 | 0.97659 | 1.192 | 3.9242
Calculations
Calibrator Flow (Qa) = 1/Slope*(SQRT(H20*(Ta/Pa))-Intercept)
Pressure Ratio (Po/Pa) = 1-Pf/Pa
% Difference = (Look Up Flow-Calibrator Flow)/Calibrator Flow*100
NOTE: Ensure calibration orifice has been certified within 12 months of use
Tisch Environmental 145 South Miami Ave, Cleves OH 45002 ● 877.263.7610 ● sales@tisch-env.com ● www.tisch-env.com
TE-6070V Calibration Worksheet